## Tabelle1
| Data | Interim Results | Standard Uncertainty ETSI in dB
Uncertainty due to Powermeter
at Level in dBm | -35 | 3.1623e-07
+
Noise in pW (with Filter No. 9) | 40 | 4e-11
Standard Uncertainty |  |  | 0.00027468
Zero Error in pW (rectangular distribution) | 50 | 0.00068673
Standard Uncertainty |  |  | 0.00039649
Calibration Uncertainty in dB (2s) | 0.07
Standard Uncertainty |  |  | 0.035
Linearity Error in % Power (2s) | 0.7
Standard Uncertainty |  |  | 0.015227
Temp. Effect 18°-28° in % Power (1s) | 0.3
Standard Uncertainty |  |  | 0.013048
Powermeter Acc. in % Power (rectangular distribution) | 0.3
Standard Uncertainty. |  |  | 0.0075287
Sum Uncertainty Powermeter |  |  | 0.041037
Uncertainty due to mismatch:
VSWR RFOUT2 of CMD | 1.5
VSWR Power Sensor | 1.1
Standard Uncertainty in dB |  |  | 0.05856
Uncertainty due to Linearity Error of FSE
Linearitiy Error FSE in dB with FSE-B22 (rect. distribution) | 0.2
Standard Uncertainty in dB |  |  | 0.11547
Uncertainty due to permissible Linearity Error of CMD in dB (rectangular distribution) | 0.3
Standard Uncertainty |  |  | 0.17321
Uncertainty of CMD level due to temperature
dB/15°K (rectangular distribution) | 0.1
Standard Uncertainty |  |  | 0.057735
Uncertainty due to ageing of CMD level
in dB per 12 months (rectangular distribution) | 0.1 |  | 0.057735
Standard Uncertainty
Combined Standard Uncertainty in dB |  |  | 0.23476
Expanded Combined Standard Uncertainty in dB with 95% Confidence (2 s) |  |  | 0.46952