R&S® UPV / UPP* Audio Analyzers Release Notes Firmware Version 4.0.5

Table of Contents

1	Information on the Current Version and History	4

1.1	Version 4.0.5.0190	4

1.2	Version 4.0.4.0159	6

1.3	Version 4.0.3.0136	9

1.4	Version 4.0.2.0107	11

1.5	Version 4.0.1.0076	12

1.6	Version 3.3.6.0834 (UPP only)	14

1.7	Version 3.3.4.0822	14

1.8	Version 3.3.2.0805	16

1.9	Version 3.3.1.0758	17

1.10	Version 3.3.0.0724	19

1.11	Version 3.2.0.0683	20

1.12	Version 3.1.2.0626 (UPV only)	22

1.13	Version 3.1.0.0571	23

2	Firmware Update	34

2.1	Update from versions 4.0.x to 4.0.5	34

2.2	Update from versions 2.x or 3.x to 4.0.5	35

2.3	Update from versions 1.x to 4.0.5	35

2.4	Local Firmware Installation	36

2.5	Remote Firmware Installation	37

3	Troubleshooting	40

4	Customer Support	43

# Information on the Current Version and History

## Version 4.0.5.0190

Attention: When updating Firmware from Version 2.x or 3.x to 4.x please carefully read section 2.2 first!

Attention: When using an UPP running Windows 7 in conjunction with the R&S Sequencing application, you will need an update of this, too.

| Component | FW version |
| --- | --- |
| UpxCheckPreconditions | 1.5.0 |
| UPVCompass_3.1.18.21-Release | 3.1.18.21 |
| Setup_4.0.50190Release | 4.0.5 |

New Functionality

Modified Functionality for all firmware versions 4.0.x

| Moving and Resizing of Panels and Graphical Windows is no longer possible with the MODIFY hard key. Use a mouse for this. | UPV |
| --- | --- |
| The functionality of the WINBAR hard key has been removed. Use the WINDOW hard key or a mouse to put the focus to a panel. | UPV |
| When a numeric input field is selected, the input mode can be switched between edit mode and variation mode with a Rotary Knob click. The enter key on a normal keyboard just confirms the actual value. | UPV |
| Navigation in menu Combo -> cursor keys |  |
| On an external keyboard ALT+F4 closes the Help window | UPV UPP |

Fixed Issues

| Limit curves changing when loaded from file and performing external sweep (794576) | UPV UPP |
| --- | --- |
| 627232 peak measurement values are incorrect when all additional functions are switched off | UPP |
| Context menu for measurement result configuration sometimes not working (586498) | UPV UPP |
| Reference signal for chirp based phase measurements is Analyzer1 signal instead of generator signal (554215) | UPV |
| PESQ 861.2 MOS-LQOW is constantly zero in PESQ/POLQA Graph 1 and 2 (524432) | UPV |
| Function Monitor = Peak with function bandpass shows measurement value by 6 dB too low (505971) | UPV UPP |
| Firmware crashes when reading status registers by K1 application (502058) | UPV |
| UPV mit FW 4.0.4 requires UPV-B3 by error when loading stereo WAV files in digital generator (501122 ) | UPV |
| Firmware unable to store *.sac setup files (449779) | UPV UPP |
| *STB? results in firmware crash when issued by K1 application (440491) | UPV |
| SENSe8:PROTocol:ERRor? has wrong return value (<n> and th string) does not contain all errors d0…d9 but only the first. (378279) | UPV UPP |
| When the UPV-B42 analyzer is operated in Master Mode (Sync to Internal Clock) then data is clocked in with the wrong slope of the bit clock. Work around: set sampling delay to -9 ns. (300545) | UPV |
| Improving signal quality for down sampling rates of 128, 16, 8 (212429) | UPV |
| Setup converter unable to convert some setups when „2nd Monitor“ is set to „Funct Monitor“ (203350) | UPV UPP |
| DC coupling not activated when setup with channel 2 measurement only is loaded (51733) | UPV |
| POLQA measurement: Ref and Deg Speech Ratio: The SCPI query commands return Deg instead of Ref, and Ref instead of Deg. (139999) | UPV |
| Cascading: M + S1 are connecting though S1 has no clock (139996) | UPP |
| Cascading: A S1 device is already registered but a S2 device gets no connection. (140015) | UPP |
| Software crash with two UPV-B48 and “Analog 16 Chan” nd Multi Channel selected (140003) | UPV |
| Cascading: problems with short term connection losses of LAN (139998) | UPP |
| User data buffers operated by SCPI get suddenly deleted for unknown reasons, e.g. when using INPut2:COUPling DC (195043) | UPV |

Known Issues

| Problems with UPV-B42 and ALaw / MLaw coding in conjunction with 8 audio bits: (818817) With slot length 8, only 8 bits are coded instead of 13 coding has some inaccuracies which may cause level errors below 150mFS | UPV |
| --- | --- |
| When switching off the sweep while fast sweep measurement is running, some panels may be updated wrong. To work around this, stop the sweep before changing anything or select the generator function a second time. | UPP |
| On some devices, UpxCheckPreconditions always showing message about “R&S Framework” deinstallation (section 2.4.1) | UPV UPP |
| Update installation may be denied with “missing .NET 2.0” framework (section 2.4.2) | UPV |
| Firmware may not work on an FMR6 after update to version 4.x.(section 2.4.3) | UPV |

## Version 4.0.4.0159

Attention: When updating Firmware from Version 2.x or 3.x to 4.x please carefully read section 2.2 first!

Attention: When using an UPP running Windows 7 in conjunction with the R&S Sequencing application, you will need an update of this, too.

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.4.3 |
| UPVCompass_3.1.18.21-Release | 3.1.18.21 |
| Setup_4.0.40159Release | 4.0.4 |

New Functionality

| THD now allows measurement beyond the first 8 harmonics. 2 new Meas Modes have been added that now allow to calculate THD result up to d50 (49th harmonic). Meas Mode Note: Only the first 8 harmonics - and the fundamental - are always displayed in the graphic irrespective of the selected measurement mode. dmin to dmax: The distortion of all harmonics dmin to dmax – root from the squared sum of these harmonics – is measured relative to the total RMS. 2 additional menu lines occur and allows the setting of the range Lev dmin to dmax: The distortion voltage of all harmonics dmin to dmax (root from the squared sum of these harmonics) is measured 2 additional menu lines occur and allows the setting of the range By setting dMin and dMax to the same value this single harmonic is calculated. d Min enter the first harmonic of the harmonic range to be calculated. Can be any number from 2 to 50 d Max enter the last harmonic of the harmonic range to be calculated. Can be any number from d Min to 50 | UPV UPP |
| --- | --- |
| Modified POLQA version 2.4 including implementers guide 2016-01 no 4 (application of the reverberation indicator removed in the SWB mode of P.863) (170261) | UPV |
| POLQA, PESQ, PEAQ: To perform offline measurements internal generator need not to be active but can be turned off via 'Channel Off' or 'Output Off' | UPV |

Modified Functionality

| Graphics profile editor now provides more convenient handling of definition files: A Click on the “Profile…” button now opens a popup menu: This allows a more clearly way for dealing with profile file definitions: When the editor is opened first time, no file is selected When changes are made and the dialog is closed a message will appear asking if the changes should be stored. Wen clicking Yes, a file selector will appear When a file definition is loaded by “Load profile…”, after asking for storing, changes always will be automatically stored in this file When “Save profile as…” is selected, the current definition may be stored with another file name When “New profile” is selected, unsaved changes will be (after asking for store) saved. Afterwards, all settings are reset to default and the editor will go back to the state with no file selected. | UPV UPP |
| --- | --- |

Fixed Issues

| Missing texts in status bar after start up / missing hints in status bar (189864, 192853) | UPV UPP |
| --- | --- |
| Firmware version crashes when a file defined generator filter does not find the filter definition file (165247) | UPV UPP |
| Firmware version 4.0.x sometimes has problems with non-ASCII characters in file names (like German umlauts) (164666) | UPV UPP |
| Numeric input fields may run into a malfunctioned state where every digit typed in will be selected and numbers no more can be entered (161507) | UPV UPP |
| Setting Chirp Based Harmonics gets very slow during SCPI operation with larger sweep times (158335) | UPV UPP |
| Firmware 4.0.x crashes when using chirp based measurement with newly introduced measurement mode PHASE (132491) | UPV UPP |
| Firmware 4.0.x may get inoperable during SCPI operation when ERROR queue gets messages (158904) | UPV |
| SQUARE generator function occasionally using equalization file although switched off (187208) | UPV |
| Hard disk gets full while doing many POLQA measurements (147974) | UPV |
| POLQA measurement does not work correctly when given ref and play files are different in length or when given play file is longer than 40s (147861) | UPV |
| Firmware crashes after several 100 POLQA measurements (147474) | UPV |
| PEAQ offline measurement crashes with play length > 20s (132484) | UPV |

Known Issues

| When switching off the sweep while fast sweep measurement is running, some panels may be updated wrong. To work around this, stop the sweep before changing anything or select the generator function a second time. | UPP |
| --- | --- |
| On some devices, UpxCheckPreconditions always showing message about “R&S Framework” deinstallation (section 2.4.1) | UPV UPP |
| Update installation may be denied with “missing .NET 2.0” framework (section 2.4.2) | UPV |
| Firmware may not work on an FMR6 after update to version 4.x.(section 2.4.3) | UPV |

## Version 4.0.3.0136

Attention: When updating Firmware from Version 2.x or 3.x to 4.x please carefully read section 2.2 first!

Attention: When using an UPP running Windows 7 in conjunction with the R&S Sequencing application, you will need an update of this, too.

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.4.3 |
| UPVCompass_3.1.18.21-Release | 3.1.18.21 |
| Setup_4.0.30136Release | 4.0.3 |

New Functionality

| Quick start panel for beginners (see user documentation) | UPV UPP |
| --- | --- |
| Sequencing and reporting tool available for download on R&S home page | UPV UPP |
| Adding Gapless / Overlapping mode for FFT based measurement functions The functionality of the FFT mode, which was up to now not changeable and is used by FFT based measurement functions, is extended to additional modes. If set to Standard, measurement is done same as before. If set to Gapless, measurements are done without loss of samples. This mode is only available with analogue analyser fix range and Meas Time Precision setting If set to Overlapping, measurements are done without loss of samples and with overlapping FFT calculations. The range of overlapping can be chosen. Running a gapless or overlapping FFT requires: - Analogue analyzer set to Fix Range - Frequency measurement time set to Precision If set to Normal, FFT measurements have some gap between subsequent acquisitions (same as before). | UPP |
| Input Monitor now more flexible: The functionality of the Input Monitor, which was up to now limited to Input Peak, is extended to a second Level Monitor. The test point of this monitor can now be set to Input, Level (after Pre filter) or Function (after function filters). UPP only: "With FFT Mode Overlap or Gapless the test point is limited to Function for both monitors." If set to Input the Input Peak only can be displayed (same as before). If set to Level or Function it is possible to display RMS, DC or PEAK. This allows displaying two different levels (chosen from RMS, DC or PEAK) in addition to the Function Result (e.g. SINAD) and Frequency. The test point Function is available only for measurement function FFT or if Post-FFT is activated, and RMS and DC are calculated in frequency domain (based on FFT bins) then. RMS and DC are always calculated in time domain for test point Level. | UPV UPP |

Modified Functionality

Fixed Issues

| Firmware crashes if loading a setup with enabled limit curves (122135) | UPV UPP |
| --- | --- |
| File selector for saving files does not take the correct path (119342) | UPV UPP |
| If an error occurs during manual operation, the SCPI error queue gets an entry (112749) | UPV UPP |
| Cursor set to min / max not working (106929) | UPV UPP |
| Number formatting problems with some settings | UPV UPP |
| Loading UPV setups does not give an error if settings of display channel 2 are different against channel 1 (116305,116289) | UPP |

Known Issues

| When switching off the sweep while fast sweep measurement is running, some panels may be updated wrong. To work around this, stop the sweep before changing anything or select the generator function a second time. | UPP |
| --- | --- |
| Measurement PEAQ Offline crashes with file length > 20s (issue concerning all releases which support PEAQ) | UPV |
| Storing of WAV files recorded with PESQ, POLQA or PEAQ does not work if Play Mode AUTOMUTE is active (issue also concerning releases 4.0.2, 4.0.1, 3.3.4, 3.3.2) | UPV |
| Generator function PLAY and Analyzer function RECORD do not start synchronously (issue also concerning releases 4.0.2, 4.0.1, 3.3.4,3.3.2, 3.3.1) | UPV |
| Chirp Based Measurement crashes if new measurement mode phase is active and generator function is changed away from Chirp and back. | UPV UPP |
| On some devices, UpxCheckPreconditions always showing message about “R&S Framework” deinstallation (section 2.4.1) | UPV UPP |
| Update installation may be denied with “missing .NET 2.0” framework (section 2.4.2) | UPV |
| Firmware may not work on an FMR6 after update to version 4.x.(section 2.4.3) | UPV |

## Version 4.0.2.0107

Attention: When updating Firmware from Version 2.x or 3.x to 4.x please carefully read section 2.2 first!

Attention: When using an UPP running Windows 7 in conjunction with the R&S Sequencing application, you will need an update of this, too.

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.4.2 |
| UPVCompass_3.1.18.21-Release | 3.1.18.21 |
| Setup_4.0.20107Release | 4.0.2 |

New Functionality

| Now fully supporting the Windows 7 operating system | UPV UPP |
| --- | --- |
| Chirp based Measurements added | UPV UPP |
| Numeric values can be modified via mouse wheel | UPV UPP |
| POLQA 2.4 (ITU Version) implemented | UPV |

Modified Functionality

| Device Key is now displayed with checksum (Help -> About / Install Options) | UPV UPP |
| --- | --- |

Fixed Issues

| Sometimes panel layout gets corrupted so that input elements cannot be used (56408) | UPV UPP |
| --- | --- |
| Inconsistent file location of graphic profiles. Must be on D:\UPV\user (56712) | UPV UPP |
| Trace reference files are not loaded automatically (55175) | UPV UPP |

Known Issues

| UpxCheckPrreconditions always showing message about “R&S Framework” deinstallation (section 2.4.1) | UPV UPP |
| --- | --- |
| Update installation may be denied with “missing .NET 2.0” framework (section 2.4.2) | UPV |
| Firmware may not work on an FMR6 after update to version 4.x.(section 2.4.3) | UPV |

## Version 4.0.1.0076

Attention: When updating Firmware from Version 2.x or 3.x to 4.x please carefully read section 2.2 first!

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.4.2 |
| UPVCompass_3.1.18.21-Release | 3.1.18.21 |
| Setup_4.0.10076Release | 4.0.1 |

New Functionality

| Now supporting the Windows 7 operating system | UPV |
| --- | --- |
| Chirp based Measurements added | UPV |
| Numeric values can be modified via mouse wheel | UPV |

Modified Functionality

| “Modify” hard key has changed meaning (section 2.2.1) | UPV |
| --- | --- |
| Keyboard Enter does not change editor into variation mode (section 2.2.1) | UPV |
| Device Key is now displayed with checksum (Help -> About / Install Options) | UPV |

Fixed Issues

| Sometimes panel layout gets corrupted so that input elements cannot be used (56408) | UPV |
| --- | --- |
| Inconsistent file location of graphic profiles. Must be on D:\UPV\user (56712) | UPV |
| Trace reference files are not loaded automatically (55175) | UPV |
| Converting of setups from UPP to UPV sometimes fails (64689) | UPV |

Known Issues

| UpxCheckPrreconditions always showing message about “R&S Framework” deinstallation (section 2.4.1) | UPV |
| --- | --- |
| Update installation may be denied with “missing .NET 2.0” framework (section 2.4.2) | UPV |
| Firmware may not work on an FMR6 after update to version 4.x.(section 2.4.3) | UPV |

## Version 3.3.6.0834 (UPP only)

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.3.3 |
| UPVCompass_2.1.77.0-Release | 2.1.77.0 |
| Setup_3.3.60384Release | 3.3.6 |

New Functionality

Modified Functionality

Fixed Issues

| HDMI E-EDID: The transmitted checksum was wrong, which caused a connected generator to stop output data transmission. | UPP |
| --- | --- |

Known Issues

## Version 3.3.4.0822

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.3.3 |
| UPVCompass_2.1.77.0-Release | 2.1.77.0 |
| Setup_3.3.40822Release | 3.3.4 |

New Functionality

| Muting after measurement, selectable for generator function Random, Play and Play+Anlr | UPV |
| --- | --- |
| UPV-K421: supports clock frequencies down to 1 kHz and sampling rates up to 400 kHz; supports down sampling factor 8; sample frequency can be set manually, if Clock Source Extern is selected | UPV |
| Averaging added for POLQA measurements | UPV |
| Undersampling and "triggered" now available for averaged FFT | UPV |
| SCPI only command: SYSTem:SEArch? 'Drive+Path+Filename' | UPV |
| If UPV-K9 is installed then all extended functionality (K6) is available | UPV |
| Generator Bandwidth selection "Sine 185 kHz" forces low distortion | UPV |
| Store trace as "Equalization List" or "Sweep List" available for 1/n octave bar graph display | UPV UPP |
| Generator phase accuracy increased for frequencies below 1 kHz | UPV UPP |
| Support added for UPP-B8 option | UPP |

Modified Functionality

| UPV-K421: clipping instead of rolling over, if 1.0 FS is exceeded | UPV |
| --- | --- |

Fixed Issues

| POLQA NB recordings can now be stored (11092) | UPV |
| --- | --- |
| Improved stability of external sweep with fix RMS and S/N measurements with precision frequency setting (11528, 11157, 11529) | UPV |
| UPV-K1: improved stability in communication with firmware | UPV UPP |

Known Issues

## Version 3.3.2.0805

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.3.3 |
| UPVCompass_2.1.77.0-Release | 2.1.77.0 |
| Setup_3.3.2.0805 Release | 3.3.2 |

New Functionality

| Allowing changes of sweep parameters frequency or volume only when sine signal has the next zero transition | UPP |
| --- | --- |

Modified Functionality

| Universal generator function: increasing maximum output level with equalization | UPP |
| --- | --- |
| FFT averaging now allows level trigger for each accumulated measurement | UPP |

Fixed Issues

| File menu --> Exit crashing sometimes (10766) | UPP |
| --- | --- |
| Digital analyzer not working after start up if generator was analogue (11310) | UPP |
| SCPI query for undefined parameters resulting in Visa Timeout (9922) | UPP |
| Wrong DC offset value measurements on channel pairs 1/2, 3/4, 5/6, 7/8 when using different AC coupling settings (10764) | UPP |
| HDMI fixed ARC problems with HDMI 1.4 CEC handshake (11553) | UPP |

Known Issues

## Version 3.3.1.0758

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.3.3 |
| UPVCompass_2.1.77.0-Release | 2.1.77.0 |
| Setup_3.3.10758Release.exe | 3.3.1 |

New Functionality

| Introducing support for R&S Sequencing and Reporting tool Can be invoked from UPP menu (File ( UPx Sequencing …) or tool bar | UPP |
| --- | --- |
| HDMI: protocol generator now setting “Channel Number” to “Don’t Care” LipSync measurement now enabled with “Sine Burst” signals | UPP |
| SCPI: hardcopy now creating missing directories HCOPy <string> allows direct output of a graphical window with its window title even if it has no focus (only with “SCPI update On”) DISPlay:SCReen <n> allows selecting of given screen (only with “SCPI update On”) | UPP UPV |
| All generated trace files now containing description of current device, serial no, date and time of creation | UPP UPV |

Modified Functionality

| HDMI: several improvements Increased performance by reducing internal initializations Using “Coding Autodetect” now rejecting trigger source “Gen Burst” Allowing “Hotplug Detect” via new button “Resync” Now ignoring HDMI interrupts while input is not Sink / ARC Reducing analyzer update rate of protocol data and video timings during continuous measurements and sweeps | UPP |
| --- | --- |
| Time Tick sweeps: no longer restarting upon sample rate or coding format changes at digital inputs | UPP UPV |
| Improved overall stability in Remote operation, especially in “SCPI update On” mode | UPP UPV |
| PESQ / POLQA improvements Degraded scaling now independent from data source and meas mode Degraded scaling now independent from creation process (store in DUT mode / Record function) Reducing gain headroom of filter usage, overall reduction up to 0dB when no filter selected | UPV |
| Graphic Display for Time Tick sweep: The x-axis can now be scaled up to 2048 ks (23.7 days). This allows to display the maximum sweep time of 1024 points with steps of 2000 s. | UPP UPV |
| External Sweep: Improved trigger delay: Up to now the trigger delay was related to the measurement time, which was specified in the analyzer measurement function, and which could be considerably long. Now the trigger has an independent short delay. | UPP UPV |

Fixed Issues

| fixed issues with HDMI: fixed initialization problem when color depth at AUX IN changes fixed “Format Change” error when CTS / N changes, no longer changing hardware settings When DSP reports a PCM signal, now using direct data routing Changing from “Monochrome” to “Pattern” or vice versa does now correctly update the audio settings | UPP |
| --- | --- |
| fixed issue with wrong SCPI status events on generator state | UPP UPV |
| fixed issue with creating hardcopies in SCPI operation:*WAI working correct | UPP UPV |
| fixed Gentrack Reference issues: Play, Play+Anlr, Peak functions and different level settings in generator chan 1 & 2 | UPP UPV |
| fixed problem with marker drawing upon small FFT sizes: markers now at correct x position | UPP UPV |
| fixed problems resulting from loading a multiscan trace file | UPP UPV |
| fixed problem with binary data transfer and UPV-K1 | UPV |
| fixed several problems with POLQA Correct labelling of graphical y axis Repainting of graphical display now showing POLQ result instead of last PESQ result Scaling of x axis now updated with “Record length” Average Delay display now allowing negative values | UPV |

Known Issues

| The printed and online documentation contains some wrong information regarding the POLQA measurement: There are only two different bandwidths, namely “narrow band” and “super-wide band” The measurements are done according to ITU-T P.863 instead of the mentioned P.861 / P.862 In narrow band mode, the POLQA measurement only supports reference files with 8 kHz sample rate | UPV |
| --- | --- |

## Version 3.3.0.0724

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.3.2 |
| UPVCompass_2.1.77.0-Release | 2.1.77.0 |
| Setup_3.3.00724Release | 3.3.0 |

New Functionality

| Supporting UPP-B4 HDMI / digital audio measurements | UPP |
| --- | --- |
| In the "Start Condition" section of the analyzer config panel the selection of “Max FFT Size” was added for limiting the automatic zooming of FFT size. Especially useful for external sweeps to reduce measurement time of FFT based functions. | UPP UPV |
| Dual domain signal generation: If the generator instrument is switched between analog and digital, the previously selected generator keeps running with all its parameters settings made before. When the sampling frequency of the digital generator is below 200 kHz, this feature is possible for all analog and digital generator settings. When the sampling frequency of the digital generator exceeds 200 kHz, the analog generator must be set to function Sine and Low Dist must be set to ON (UPV-B1 option required). | UPV |

Modified Functionality

| Improvement of setup and SCPI compatibility between UPV and UPP | UPP |
| --- | --- |
| Speedup of external sweep: start trigger level now 0.1% below nominal value | UPP UPV |

Fixed Issues

| Loading of sweep list or equalization files crashed when > 1024 points. Now loading is rejected. | UPP UPV |
| --- | --- |
| External sweep sometimes stopping at 2nd point | UPP UPV |
| POLQA in offline mode crashed when example WAV file was selected | UPV |

Known Issues

| The remote control commands MMEM:MSIS and MMEM:CDIR, which could be used to define a default folder for file storage, do not work properly. The default folder is always c:\upv\bin and cannot be changed with those commands. Work-around: Specify full path names in commands for storing and loading files, e. g. MMEM:STOR:STAT, “d:\upv\user\test1.set” | UPV UPP |
| --- | --- |

## Version 3.2.0.0683

Firmware package contents

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.3.2 |
| UPVCompass_2.1.77.0-Release | 2.1.77.0 |
| Setup_3.2.00683Release | 3.2.0 |

New Functionality

| Perceptual Objective Listening Quality Analysis POLQA according to ITU-T recommendation P.863 is now available with software option UPV-K63. | UPV |
| --- | --- |
| Digital Analyzer for Pulse Density Modulated data streams is now available with software option UPV-K421 as add-on to UPV-B42 hardware option. | UPV |
| Measurement functions PESQ, PEAQ, POLQA now available with digital and analogue measurements. Reference file sample rate must be multiple of 8 kHz | UPV |
| Remote control emulation mode for direct replacement of the obsolete HP audio analyzer 8903B in remote controlled system applications. See also Application Note 1GA54_2e. | UPV UPP |
| Analyzer function THD: The automatically optimized FFT size can be increased with a “refinement factor” (2, 4, or 8) in order to reduce noise contributions to the calculation of the harmonics. Especially useful for noisy signals without significant harmonics. | UPV UPP |
| New “factory preset” functionality added during start up of firmware when the “OUTPUT OFF” indicator (UPV) or “BUSY” indicator (UPP) is flashing for 2 seconds: Pressing PRESET key (UPV) or TRIGGER key (UPP) during this period of time will delete the current settings and reset the display resolution to 800x600 after confirmation with the same key. | UPV UPP |
| Generator function Stereo Sine: Filter and Equalizer can be independently selected for each channel. Not yet available via remote control. | UPV UPP |
| Generator function Play: Filter can be independently selected for each channel. Not yet available via remote control. | UPV UPP |
| Protocol analysis in line with IEC 60958 now supported with software option UPP-K21. Digital Audio option UPP-B2 required. | UPP |
| 1/n octave analysis now available with software option UPP-K601. | UPP |
| 600  output impedance (switchable) and 600  input impedance (switchable) is now available for instruments with serial numbers 100100 or higher. This feature is not retrofittable into older instruments. | UPP |
| Monitor outputs now have 0  output impedance an can provide a DC voltage for instruments with serial numbers 100100 or higher. 0  output impedance is not retrofittable into older instruments. | UPP |

Modified Functionality

None

Fixed Issues

| Analyzer functions PESQ, PEAQ, POLQA: From now on the measurement terminates with a fail message as soon as an overrange is detected. | UPV |
| --- | --- |
| Analyzer functions FFT or 1/n octave analysis might not terminate when measured on input channel 2 only. | UPV UPP |
| Sweep graphics: If the sweep steppes down from higher to lower values (i. e. the curve is plotted from right to left) the START button did not clear the scans. Now correct. | UPV UPP |
| Deadlock might occur when executing *wai or *opc? | UPV |
| Generator function Chirp: Now restarted on start of measurement. | UPV |

Known Issues

| The remote control commands MMEM:MSIS and MMEM:CDIR, which could be used to define a default folder for file storage, do not work properly. The default folder is always c:\upv\bin and cannot be changed with those commands. Work-around: Specify full path names in commands for storing and loading files, e. g. MMEM:STOR:STAT, “d:\upv\user\test1.set” | UPV UPP |
| --- | --- |

## Version 3.1.2.0626 (UPV only)

Firmware package contents

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.3.2 |
| UPVCompass_2.1.77.0-Release | 2.1.77.0 |
| Setup_3.1.20626Release | 3.1.2 |

New Functionality

| Analyzer/Generator |  |
| --- | --- |
| Analyzer function THD: If all harmonics are outside the upper bandlimit, then the THD result is now set to “invalid”. Up to now the result was displayed as 0 % or -240 dB. | UPV |
| Analyzer functions THD and THD+N: Undersampling FFT can now be activated to allow high-resolution, low-frequency distortion measurements. Post-FFT must be enabled, FFT size must be set manually. | UPV |
| Analog analyzer: Automatic internal DC offset correction now implemented via setting “Anl.ZeroAuto = ON” in diagnostic panel | UPV |

Modified Functionality

None

Fixed Issues

| Analyzer function FFT: Equalization did not work, though activated and a valid file was loaded. Now correct. | UPV |
| --- | --- |
| When a bandpass is used for analyzer or generator, it may become instable, if: 1. upper passband edge >> lower passband edge (very broad passband), 2. low signal frequency 3. high signal level (> 0.5 of FS) Now correct. | UPV |
| Digital Audio Interface (UPV-B2): - Sync Output is always set to Internal clock, - Sync Output Type is always set to Wordclock, independent of the settings made in the analyzer config panel. Now correct. | UPV |
| If the generator is overloaded, it will be switched off internally. But this state is not indicated by the Overload LED, nor by the Output Off LED, nor by the Gen Ovld status information. Now correct. | UPV |
| Multichannel Analyzer: If a new channel was selected in the display panel, the trace of the previous channel was not removed. Now correct. | UPV |
| Analog analyzer: 250 kHz bandwidth: special combinations of filters together with very long measurement times might cause the measurement not to terminate. Fixed by reducing the maximum measurement times. | UPV |
| Generator overload might sporadically occur, if UPV-B3 is installed and generator bandwidth is changed at high output signal levels. Fixed by modification of hardware setting procedure. | UPV |
| Phase measurement did not work with dynamic mode precision (distortion measurements). | UPV |
| Generator level sweep: if start and stop level are defined in dBm, the x-axis of the sweep graph panel is scaled to the open circuit level of the generator – instead of the settings defined in the generator function panel. Now correct. | UPV |
| Waveform display might be corrupted if: - continuously measuring RMS in auto fast mode - signal frequency below 200 Hz - waveform length > 10 ms | UPV |

Known Issues

| The remote control commands MMEM:MSIS and MMEM:CDIR, which could be used to define a default folder for file storage, do not work properly. The default folder is always c:\upv\bin and cannot be changed with those commands. Work-around: Specify full path names in commands for storing and loading files, e. g. MMEM:STOR:STAT, “d:\upv\user\test1.set” | UPV UPP |
| --- | --- |
| Graphics: The multichannel mode is vailable for FFT and bargraph window. Though it should be able to display more than 2 channels per graphics window, only 2 channels are possible now. | UPV UPP |

## Version 3.1.0.0571

Firmware package contents

| Component | FW version |
| --- | --- |
| UPxCheckPreconditions | 1.3.1 |
| UPVCompass_2.1.77.0-Release | 2.1.77.0 |
| Setup_3.1.00571Release | 3.1.0 |

New Functionality

| Analyzer/Generator |  |
| --- | --- |
| Analyzer functions FFT, THD and THD+N: A progress information is displayed in the status bar, if averaging is selected and if the averaging takes more than only a few seconds. The progress information starts after the 2nd average step and is updating once per second. | UPV UPP |
| PESQ now available in digital domain for I2S and USI Dual Channel instruments | UPV |
| New features for generator function Play: Restart: Off: the waveform is not restarted Auto: the waveform is restarted after any user interaction Once: the waveform is started once If restart is switched from Auto to Off, the generator is stopped and muted. Auto is the default to be compatible to older firmware versions Remote control commands: SOURce:PLAY:RESTart OFF SOURce:PLAY:RESTart AUTO SOURce:PLAY:RESTart ONCE Inter-channel Delay Available only for stereo wav files, played in stereo mode. Maximum delay is 48 k samples (1 s @ 48 kHz sampling rate) Remote control commands: SOURce:PLAY:DELay[1] .25 S SOURce:PLAY:DELay2 0 S Stereo level setting The level can now be set independently for each channel. For mono and stereo wavefiles. Two voltage setting modes are supported (same as for stereo sine) Channel 1 and ratio Ch2/1 This is the default, with ratio = 1 (mono level setting, thus compatible to previous firmware versions) Remote control commands: SOURce:VOLTage:SELect VLRT SOURce:VOLTage 0.25 V SOURce:VOLTage:RATio 0.5 Ch1&2 voltage setting Remote control commands: SOURce:VOLTage:SELect VLVL SOURce:VOLTage 0.15 V SOURce:VOLTage:CH2Stereo 0.30 V | UPV |
| Generator Sweep: The new menu item Halt defines the frequency of the generator after the sweep has terminated. Start:The generator returns to the start frequency Value: The generator frequency is set to a predefined value Mute: The generator mutes after the sweep has terminated. Remote control commands: SOURce:SWEep:FREQuency:HALT STARt SOURce:SWEep:FREQuency:HALT VALue;:SOURce:SWEep:FREQuency:HALT:VALue 1 KHZ SOURce:SWEep:FREQuency:HALT MUTE This halt state can be defined for all sweepable parameters. Therefore instead of FREQency also VOLTage, PHASe (for stereo sine only), INTerval and ONTime (for Burst only) is supported. | UPV UPP |
| Play before measurement: This feature is only available for the 2-channel analog analyzer, with UPV-K9 option installed. It can then be activated in the analyzer config panel: It is intended for use in mobile phone tests, to play an activation signal (usually a voice-like signal stored in a *.wav file) before the test signal (e.g. a sinewave) is started. Measurement starts, with the start of the test signal, and can be further delayed with the start condition delay. As test signal one of the generator functions must be selected. An appropriate measurement function must be selected in the analyzer. The *.wav file must have been loaded in the generator function play, before the test signal function is selected. Remote control command: TRIGger:PLAYbefmeas ON ¦ OFF |  |
| UPV-B42 generator: the maximum setting of MClk Ratio is increased from 768 to 800. | UPV |
| The THD+N&SINAD measurement function now provides 2 filters. Up to now only one filter was possible. | UPV UPP |
| Measurement functions THD+N&SINAD and THD: Average Mode for the Post FFT The number of averages can be set in the range from 1 to 10000 The resulting measurement value is calculated from the averaged FFT. Remote control commands: SENSe:FUNCtion:FFT:AVERage:MODE NORMal SENSe:FUNCtion:FFT:AVERage 12 | UPV UPP |
| Analyzer function record Up to now only the automatic mode was supported. From now on 3 fixed settings are possible. Remote control command: SENSe[1]:FUNCtion:RECord:BPS AUTO ¦ L8 ¦ L16 ¦ L32 | UPV |
| Interchannel delay and impulse response can now be calculated with the analyzer function Transfer&Co from any input signal, maybe speech, music or noise. | UPV |
| Noise density can now also be displayed in the unit "dBV/√Hz". The reference must be set to "Noise Density dB" in the FFT graph config panel. | UPV |
| Analyzer function FFT: The RMS level via FFT can now be calculated within a freely selectable frequency range Remote control command: SENSe[1]:FREQuency:LIMit ON ¦ OFF OFF is the same as the previously supported wideband mode (default). | UPV UPP |
| Analyzer function FFT: The start of the FFT can now be triggered simultaneously with the waveform trigger. The FFT triggered mode is activated in the analyzer FFT function . If not yet active, the waveform is switched on. The trigger parameters are the same as for the waveform function, and hve to be defined there. | UPV |
| Remote Control |  |
| New remote control command: INITiate:CONTinuous RSTart This commands starts a single measurement (like INITiate:CONTinuous OFF), and in addition: - resets all parameters (like Min/Max values or limit violations) - re-starts the wavefile for the generator function Play Thus this command simulates the sequence of pressing the START and then the SINGLE button on the front panel. | UPV UPP |
| New remote control command: SYSTem:WINStyle ON ¦ OFF | UPV UPP |
| New remote control commands: SENSe[1]:CONFig:COPYother ONCE SENSe2:CONFig:COPYother ONCE SENSe6:CONFig:COPYother ONCE These commands copy the settings made for channel 1 in the Display Config Panels of Function, Input and Level Monitor to channel 2. | UPV UPP |
| Strings can be transmitted to and queried from the UPx, by using the remote control commands SYSTem:MEMory:STRing<i> 'Any String' and SYSTem:MEMory:STRing<i>?. If the queried string was an empty string ‘’ this could lead to a timeout in the controlling program. If the string buffer was deleted before, the query answer was ‘ ‘ (space). Now in both cases the query answer is ’empty’. | UPV UPP |
| New command to query the MAC address of the instrument: SYST:SINFo:MAC? | UPV UPP |
| New remote control command… SENSe<n>:DATA:ALL? … to query the measurement results of all channels with one command. In the return string, the results a separated by commas. | UPV UPP |
| New remote control command… OUTPut:IMPedance:UNBalanced? … to query the value (which is fixed to 5 ohm) of the generator impedance, when the generator is set to Output Type Unbal. | UPV UPP |
| Graphics |  |
| Grafic Windows can have a title above the grid. The title is defined in the graphic config panel: The title of a graphic is stored (and loaded) with the trace file. | UPV UPP |
| Numeric display Up to now the channel names were fixed to Ch1, Ch2, …….. , Ch16 Now the names for the channels can be freely configured in the config Panel, for example: ….. for the following display….. Remote control command: SYSTem:CHNString 'Left;;Right;;Center;;LFE’ | UPV UPP |
| The numeric display can now be copied to file or clipboard. Up to now this was only supported for active graphics window. | UPV UPP |
| General |  |
| The setup of the instrument can be stored as *.html file, by selecting “Report Files” as file type when saving the setup….. …..and loaded with a browser | UPV UPP |
| Trigger Input now implemented The impuls at the trigger input acts like a keystroke of one of the front panel keys START, SINGLE or STOP/CONT. The Mode selects, which key is simulated by the trigger impulse. Edge selects the active slope of the impulse. It is expected, that the impulse has a fast rise/fall time due to modern logic circuit outputs. Trigger input cannot be used together with a generator or analyzer sweep. Remote control commands: AUXiliaries:TRIGger:INPut:ENABle ON ¦ OFF AUXiliaries:TRIGger:INPut:MODE MSTart ¦ MSINgle ¦ MCONtstop AUXiliaries:TRIGger:INPut:EDGE RISing ¦ FALLing | UPV |

Modified Functionality

None

Fixed Issues

| Analyzer functions PESQ and PEAQ: When operating in offline mode, where the degraded signal is loaded from file, the waveform display still showed the signal at the analyzer input connectors, which is meaningless in offline mode. Now the signal in the loaded file is displayed. | UPV |
| --- | --- |
| Analog analyzer bandwidth 250 kHz: If a prefilter and 3 function filters (if possible for the selected measurement function) were used, the measurement did not terminate. From now on the third function filter can no longer be selected for the 250 kHz bandwidth analyzer. | UPV |
| Generator signal sine squared burst: If a negative voltage was entered the generated pulse was still positive and the level setting was not correct. Now the pulse is inverted by entering a negative voltage. | UPV |
| Sweep Graph Config panel: If Inp Peak was selected as Y-Source, the unit could not be selected manually, though Unit Fnct Track was unchecked. Now correct. | UPV |
| Analog analyzer with bandwidth 250 kHz: If a startcondition delay > 2.8 s was used, the user interface was blocked (due to internal calculation errors). The UPV application had to be ended with the windows task manager. Now fixed. | UPV |
| Multichannel instruments: Wrong input peak reading in the 3rd and 4th active channel, if the measurement time was longer than 500 ms. Now correct. | UPV |
| Waveform trigger channel, analog analyzer bandwidth 250 kHz, single channel measurement: When the channels were switched between 1 and 2, the trigger source was set to “Manual”. Now the trigger source tracks to the selected channel. | UPV |
| Analog 16 channel instrument (2 UPV-B48 options fitted): Channels 9 to 16 could not be set to a fixed range via remote control, because the command SENSe:VOLTage:RANGe<n>:VALue <level> produced a "Data out of range" error, if <n> was in the range from 9 to 16. Now correct. | UPV |
| Generator function Stereo Sine: The input range of the voltage ratio setting was not limited according to the maximum allowed level. Now correct. | UPV |
| Generator function Play: The level of channel 2 was too low, when channels were switched between Stereo and Mono and if the level was > 50 % of Fullscale. Now correct. | UPV |
| Up to now the limit file parser expected capital letters. An error message occured, when e.g. Hz was written instead of HZ. Now the parser is no longer case-sensitive. | UPV |
| The following remote control commands to show/hide and define the legend string have been missing for the bargraph display: DISPlay:BARgraph1¦2:A¦B:LEGend:SHOW ON ¦ OFF DISPlay:BARgraph1¦2:A¦B:LEGend:DESCription "String" | UPV |
| The following remote control commands have been missing for the PESQ display: DISPlay:PESQ:X:LABel:AUTO ON ¦ OFF DISPlay:PESQ:X:LABel:USER 'MyXLabel' DISPlay:PESQ:X:UNIT S ¦ MIN ¦ DS DISPlay:PESQ:X:UNIT:AUTO ON ¦ OFF DISPlay:PESQ:X:UNIT:USER 'MyXUnit' | UPV |
| The following remote control commands have been missing for the PESQ display: DISPlay:PESQ1¦2:A¦B:LABel:AUTO ON ¦ OFF DISPlay:PESQ1¦2:A¦B:LABel:USER 'MyLabel' | UPV |
| The following remote control commands did not work for trace B of the PESQ display DISPlay:PESQ1¦2:B:YSOurce PEMO ¦ DELay ¦ REFSignal ¦ DEGSignal ¦ DROPouts Now correct | UPV |
| When undersample FFT was selected ….. ….. the MIN/MAX query SENSe:FUNCtion:FFT:CMPFactor? MIN ¦ MAX did not work properly: the return value was the actual value, not MIN or MAX. Now correct. | UPV |
| Multichannel instruments: The remote control command SENSe3:FREQuency:REFerence:MODE to store the frequency reference value did not work for the following parameters: MREFchannel CH<n>Store <n> = 3 to 16 | UPV |
| FFT Phase Ch1, Ch2 or Ch2-Ch1 (selected in FFT Graph Config Panel) could only be displayed, if Freq/Phase measurement in Analyzer Function Panel was set to Freq & Phase. Now fixed. | UPV |

Known Issues

| The remote control commands MMEM:MSIS and MMEM:CDIR, which could be used to define a default folder for file storage, do not work properly. The default folder is always c:\upv\bin and cannot be changed with those commands. Work-around: Specify full path names in commands for storing and loading files, e. g. MMEM:STOR:STAT, “d:\upv\user\test1.set” | UPV UPP |
| --- | --- |
| Graphics: The multichannel mode is vailable for FFT and bargraph window. Though it should be able to display more than 2 channels per graphics window, only 2 channels are possible now. | UPV UPP |

# Firmware Update

Before updating the firmware, it is recommended to save the currently installed version on C:\. The Windows Recovery and Backup Tool is described in section 2 of the operating manual.

Connect mouse and keyboard to the instrument.

Connect external monitor to UPP (DVI monitor) or UPV66 (VGA monitor).

Switch on the instrument.

Files needed for the firmware update:

UpxCheckPreconditions_1.5.0.exe

UPVCompass_3.1.18.21-Release.exe

UPxSetup_4.0.5016xRelease.exe

| ATTENTION |  | It is strongly recommended to do no firmware downgrade below the version the instrument was delivered originally (this is the version you can see if you select “Factory Default” starting with Backup/Recovery). |
| --- | --- | --- |

## Update from versions 4.0.x to 4.0.5

- Open the Windows Control Panel

- Select “Add or Remove Programs” (Windows XP) or 
“Programs and Features” (Windows 7)

- Remove old version of UpxCheckPreconditions_1.4.x

- Remove old version of UPX Firmware 4.0.x.xxx

- It is not necessary to remove other components

Continue with section 2.4 Local Firmware Installation

## Update from versions 2.x or 3.x to 4.0.5

This firmware update is a major update. All parts of the old firmware and Compass have to be removed before the new parts may be installed.

Use the Windows tool “Add or Remove Programs” to remove the R&S Framework, the R&S InstallationAssistantService, the R&S Msi_LXIWebServer and the previous version of the instrument firmware (e.g. 3.3.4.822) …

In case you missed these steps and just start the new firmware installer a prompt will occur which reminds you to remove the R&S Framework and the old firmware.

Please check if the directory “C:\Pogram Files\Rohde&Schwarz” has been deleted after this. If not, do it in the Explorer. If the directory is read only, remove the check mark in the Properties Dialog of the directory and apply the changes to all subfolder and files.

Continue with section 2.4 Local Firmware Installation

| ATTENTION |  | The UPP cannot be updated directly from a firmware version below 3.3.4 to 4.0.x. An intermediate installation of firmware version 3.3.4 is necessary. This is described in section 3.4.5. |
| --- | --- | --- |

## Update from versions 1.x to 4.0.5

For any update of UPV firmware from a version 1.x to a version higher than 2.0.0 the upgrade to version 2.0.0 is required before. For this purpose a bootable CD “UPV Audio Analyzer Firmware Upgrade 2.0.0” is provided, which is included in the “UPV Firmware Version 2.0.0 Installation Manual”, ident number 1406.0154.42.

Please contact your local R&S agency to receive the Installation Manual and the CD-ROM free of charge.

Continue with section 2.2 Update from versions 2.x or 3.x to 4.0.5

## Local Firmware Installation

The local firmware installation is performed directly on the instrument, typically using a USB memory stick which contains the files for the firmware update.

### Install the UpxCheckPreconditions

Double-click the item UPxCheckPreconditions_1.5.0.exe and start the installation by selecting "Next >" and "Install".

The UpxCheckPreconditions tool will show a dialog which confirms the device type (UPV / UPV66 or UPP) and the installed Windows version.

Click “OK”.

When updating from versions 2.x or 3.x the following prompt should appear:

The instrument is now prepared for the installation of the new R&S COMPASS Framework and the new R&S Firmware.

Click "OK" and leave the Preconditioner Installer with "Exit".

### Install all parts of the new COMPASS

Skip this section when updating from a version 4.0.x. COMPASS is already installed.

Double-click the item UPVCompass_3.1.18.21-Release.exe and start the installation by selecting "Next >" and "Install".

Skip a request for a reboot, which may be displayed on screen.

If the hard disk has less than 1GB free disk space, following errors may occur:

There are two ways to solve the problem:

- Make sure that there is more than 1 GByte free space on drive C.

- If this is not possible, reduce the needed disk space from 1000 MByte to 500 MByte. This should only be done by experienced users, because a change in the registry is necessary. Open the registry editor and change the setting of the key “HKEY_LOCAL_MACHINE\Software\Rohde&Schwarz\SoftwarePlatform\ServiceConfiguration\InstallServerEx\NeededDiskSpaceMB” to decimal 500.

When you have done this, run the COMPASS installation again.

### Install the new firmware

Double-click the item UpxSetup_4.0.xxxxxRelease.exe and start the installation by selecting "Next >" and "Install".

It may occur that the device is missing the Microsoft .NET framework 2.0 which will be needed for the new firmware version. 
Section 3.4.2 describes how to deal with this situation.

It may occur that the hard disk has less than 1GB free disk space. In this case  following errors will occur:

The solution for this problem is described in the previous section 2.4.2.

The instrument shuts down automatically.

During reboot the firmware may not start correctly due to errors as described in section 3.4.3 or 3.4.4.

Continue with section 3 Troubleshooting to fix those errors.

## Remote Firmware Installation

The firmware installation can alternatively be performed over a LAN or GPIB connection from an external PC.

  - Connect the remote instrument to your PC

  - Run  UPxSetup_4.0.5016xRelease.exe  on your PC

  - Select “Remote Installation” 
 
and click the button “Next >”

  - After scanning GPIB connections and your LAN subnet all found instruments are listed. Select the instruments you want to update:

  - If no audio analyzers are found even if they are connected to GPIB or LAN, please ensure that the Device Group is set to “AUDIOANALYZER” and perform a rescan.

  - To store Device Group “AUDIOANALYZER” permanently, select 
"Options …"-Button and type “AUDIOANALYZER” in the field shown below, 

then click "Save Options"-Button and "OK"-Button in sequence.

| ATTENTION |  | Please be careful and check twice if you have selected the correct instruments. Depending on your company’s network structure also instruments of other departments will show up! |
| --- | --- | --- |

  - Additional help will be displayed after clicking the button “Help”

  - Start installation by selecting “Install”

  - When installation is completed the instrument automatically shuts down.

  - UPV only: switch instrument off.

  - Switch instrument on again.

Depending on the installed version and the update version of the firmware the installer may request that you run the item UPxCheckPreconditions.exe before you can install the new firmware. In this case, you are not able to do the update via remote.
Please follow the instructions in section 2.4 Local Firmware Installation.

# Troubleshooting

### UpxCheckpreconditions still complaining about “R&S Framework”

This problem may occur if the uninstallation of the old COMPASS version did not remove all files completely. Section 2.2 describes how to fix this.

### Firmware update not possible (.NET 2.0 missing)

On older instruments having a Windows XP version below 03.60 following error will occur when running the firmware installation:

In this case, select “No” and break the installation.

It is recommended to download .NET Framework 3.5 from the Microsoft internet site.

https://www.microsoft.com/de-de/download/details.aspx?id=25150

After having installed this, please repeat the firmware installation.

### UPV only: Firmware does not start correctly after update

If you missed doing the steps as described in section 2.1 errors similar to following may occur during firmware startup:

Front Panel LEDs are not turned on.

Fix incorrect installation of device driver

- Open the Control panel and select “System”

- Select the “Hardware” Tab

- Open the Device Manager

- Open  the UPV Hardware node and right click on UPV Driver to open the context menu

  - Select “Update Driver”

  - Select “No, not this time”, click “Next”

  - select "Install from a list or specific location …", click “Next”

  - select "Don't search …", click “Next”

  - select the upper UPV Driver with left mouse click, click “Next”

- Shut down and restart the instrument.

### UPP only: Device does not start after update

When updating from a version 3.x on a UPP the following message will occur:

Continue with “OK“. The installation will finish successfully.

### UPP only: Update from a firmware below 3.3.4 to 4.0.x

The UPP cannot be updated directly from a firmware version below 3.3.4 to 4.0.x. An intermediate installation of firmware version 3.3.4 is necessary. This firmware consists of the files

- UPxCheckPreconditions_1.3.3.exe

- UPVCompass_2.1.77.0-Release.exe

- Setup_3.3.40822Release.exe

which can be downloaded from the R&S internet site.

- Install firmware version 3.3.4. The UPP shuts down automatically.

- Switch on UPP

- During start-up this message will appear

After this update the UPP shuts down automatically.

- Switch on UPP

- Un-install version 3.3.4 using the Windows tool “Add or Remove Programs”

- Install firmware version 4.0.4.159

Without this intermediate installation of version 3.3.4 the Flash eeprom is not updated. A firmware version 4.0.x can be installed, but the analyzer measurement functions Waveform, FFT and all FFT-based functions e.g. THD+N will not work.

# Customer Support

Technical support – where and when you need it

For quick, expert help with any Rohde & Schwarz equipment, contact one of our Customer Support Centers. A team of highly qualified engineers provides telephone support and will work with you to find a solution to your query on any aspect of the operation, programming or applications of Rohde & Schwarz equipment.

Up-to-date information and upgrades

To keep your instrument up-to-date and to be informed about new application notes related to your instrument, please send an e-mail to the Customer Support Center stating your instrument and your wish. We will take care that you will get the right information.

Regional contact

Europe, Africa, Middle East

Phone	+49 89 4129 12345

customersupport@rohde-schwarz.com

North America

Phone	1-888-TEST-RSA (1-888-837-8772)

customer.support@rsa.rohde-schwarz.com

Latin America

Phone	+1-410-910-7988

customersupport.la@rohde-schwarz.com

Asia/Pacific

Phone	+65 65 13 04 88

customersupport.asia@rohde-schwarz.com

China

Phone	+86-800-810-8228 / +86-400-650-5896

customersupport.china@rohde-schwarz.com

# Text boxes

* Namely the models UPV, UPV66, UPP200, UPP400, UPP800

© 2019 Rohde & Schwarz GmbH & Co. KG

81671 Munich, Germany

Printed in Germany – Subject to change – Data without tolerance limits is not binding.

R&S® is a registered trademark of Rohde & Schwarz GmbH & Co. KG.

Trade names are trademarks of the owners.

The following abbreviations are used throughout this document:

R&S®UPV / UPP is abbreviated as UPV / UPP.

Release Notes

Test and Measurement